## Slide 1: Advanced PAP Titrations
Karin Johnson, MD
Vice Chair and Associate Professor,
Department of Neurology, UMMS- Baystate
Medical Director, Baystate Health Regional Sleep Program
@drsleepykarin

## Slide 2: Disclosure
No Disclosures
I will be comparing protocols/functionality of different PAP devices from specific manufacturers
http://www.petpoisonhelpline.com/poison/beta-blockers/

## Slide 3: Summary
How Advanced PAP works
Titration Tips
Insurance Criteria
Titration Case

## Slide 4: Advanced PAP Settings
Mode | 
BPAP ST | Constant EPAP, IPAP, PS all night Backup rate ensures breath within time period Control inspiratory/expiratory lengths
SV | Fixed or auto EPAP in response to obstruction Variable IPAP/PS targeting recent volumes or ventilation Backup rate and inspiratory/expiratory control
VAPS | Fixed or autoEPAP in response to obstruction Variable IPAP/PS targeting set goal volume or ventilation Backup rate and inspiratory/expiratory control

## Slide 5: Inspiration Expiration Decision Circuitry
Inline sensor measuring flow or pressure drop across a restriction in the air delivery tube
Start of exhalation: Flow into lungs = 0
Inhalation- Flow(t) > average flow rate
Exhalation- Flow(t) < average flow rate
Sanders, US Patent 5,433,193
Flow | Rate of change | Fuzzy Phase
Zero | Increasing | Start Inspiration
Small positive | Increasing slowly | Early Inspiration
Large positive | Steady | Peak Inspiration
Small positive | Decreasing slowly | Late Inspiration
Zero | Decreasing fast | Start Expiration
Small Negative | Decreasing slowly | Early Expiration
Large Negative | Steady | Peak Expiration
Small Negative | Increasing slowly | Late Expiration
Zero | Steady | Expiratory phase
Berthon Jones, US Patent 8857430

## Slide 6: Respiratory Control Settings
Tigger sensitivity
  inspiratory flow above which EPAP -> IPAP
Cycle sensitivity
  inspiratory flow below which IPAP -> EPAP
Rise time
  length of time to transition from EPAP to IPAP
Inspiratory Time
  Time spent in inspiration

## Slide 7: Respiratory Control Settings
Setting | Normal | COPD | Obesity Hypoventilation | Restrictive/ Neuromuscular
Ti | 0.3 sec-2.0 sec | 0.3-1.0 sec | 0.8-2.0 sec | 0.8-1.5 sec
Rise Time | 300 ms | 150 ms | 150-300 ms | 300 ms
Trigger | Medium | Medium | Medium-High | High
Cycle | Medium | High | Low-Medium | Low
COPD - Chronic obstructive pulmonary disease; sec - second; Ti - inspiration time Adapted from ResMed Sleep Lab Titration Guide
Start with normal settings then add to see if helps comfort or breathing metrics

## Slide 8: Differences in BPAP ST
Respironics
  Spontaneous breaths are not affected by Ti settings
Resmed
  Spontaneous breaths are affected by Ti settings
  Can use respiratory control settings on both BPAP and BPAP ST

## Slide 9: Physiology affecting COPD Titrations
AutoPEEP and Dynamic Hyperinflation
  Narrowed bronchiole and decreased elastic recoil
    Slower flow in and out of lung
    More collapse on expiration when positive pressure
  More compliant alveoli
    Inspiratory flow remains higher longer especially when on BiPAP
    May require active exhalation when on BiPAP to cycle
More dead space
  Higher minute ventilation to clear CO2
Airway supported by connective tissue
Elastic Recoil
+
+
Broken alveolar attachments =
Airway collapse worse on expiration
Decreased Elastic Recoil = Slower expiratory flow
Excess mucus, hypertrophy of submucosal glands, abnormal epithelium
+
+
+
NORMAL
COPD
Notes: Broken alveolar attachments, abnormal epithelium, excess mucus, hypertrophy of submucocal glands- airway doesn’t stay open on expiration

## Slide 10: Titrating COPD patients
Look for improvement in Tidal Volumes first and then reductions in CO2
Preventing dynamic hyperinflation
  If worsening with increasing EPAP or PS then likely inducing hyperinflation
Titration change | Improved Physiology | Mechanism
Higher EPAP | AutoPEEP | Kept airways more open by offsetting AutoPEEP
High IPAP | Airway Collapse Work of Breathing | Prevent airway collapse
High Cycle And/or Shorter Ti | Decreased Elastic Recoil Work of Breathing AutoPEEP | Limit Inspiratory Flow Prolong Exhalation
Slower RR | Impaired exhalation | Prolong Exhalation

## Slide 11: Hi-PPV in COPD
Murphy JAMA 2017
BiPAP ST
Ti 0.8-1.5 RT 150-200 ms
Started 18/4 RR 14-16
Mean IPAP 24 (22-26)
BPPV targeted to goal ↓PaCO2 3-7 mmHg and SpO2 >88%
Patients with chronic COPD with frequent exacerbations
Severe COPD mean PaCO2 57
12-month risk of readmission or death
  63% in O2 +Hi-PPV
  80% in O2 alone
  ARR 17% (NNT 6)
34% reduction in acute exacerbation frequency

## Slide 12: CO2 targets by Disease
COPD: to reduce by 3-7 mmHg or 20% or <48 mmHg (Murphy 2017)
Obesity Hypoventilation: to get tidal volume over 6-8 ml/kg (NIPPV task force)
Neuromuscular: aim to keep CO2 at wake levels

## Slide 13: Advanced PAP Settings
Mode | 
BPAP ST | Constant EPAP, IPAP, PS all night Backup rate ensures breath within time period Control inspiratory/expiratory lengths
SV | Fixed or auto EPAP in response to obstruction Variable IPAP/PS targeting recent volumes or ventilation Backup rate and inspiratory/expiratory control
VAPS | Fixed or autoEPAP in response to obstruction Variable IPAP/PS targeting set goal volume or ventilation Backup rate and inspiratory/expiratory control

## Slide 14: Problem with CPAP/BPAP/BPAP ST
Breathing patterns/ ventilation that change throughout the night
Not enough pressure to ventilate in REM and supine
Too high pressures in NREM or lateral position inducing
Central events
PAP intolerance
Mask leak
Hyperinflation in COPD
Breathing patterns that change breath to breath
  Periodic breathing
Too high pressure drops CO2 levels and increases breathing instability triggering periodic breathing
Not enough pressure to prevent obstruction

## Slide 15: Tip 1: Watch for BPAP ST lengthening Rhythmic Central apneas
BPAP
BPAP ST
Longer central apneas
Flow is backup rate kicking in
More instability/arousals

## Slide 16: Tip 2: Watch for Oxygen Banding Pattern
Sign of periodic breathing
  Rhythmic desaturations to similar degree
  Rarely below 80%
  Improves in REM

## Slide 17: Tip 3: Score Central Hypopneas
30% reduction in pressure transducer for 10 seconds with arousal or 3% O2 desat AND NO
  Snoring OR
  Flow limitation OR
  Paradoxical breathing

## Slide 18: Tip 4: Watch for Mixed Events
Decreased drive often causes airway collapse
BPAP may can squaring of flow signal
Think central physiology even if a little snoring and flow limitation if
  Worse NREM/Better REM
  Symmetric periodic waxing and waning of both flow and effort
  Worsening with increasing PAP pressures
  Arousals at height of respiratory effort

## Slide 19: Tip 5: Recognize Different Types of Central Physiology
Rhythmic
  High loop gain, driven by CO2 fluctuations
  Backup rate can worsen
Irregular
  Driven by central breathing suppression (narcotics)
  Neuromuscular
  Backup rate can help

## Slide 20: Servoventilation (SV)
Pressure Profile | Goal | Uses/ Cautions
Fixed or auto EPAP in response to obstruction Variable IPAP/PS targeting recent volumes or ventilation Backup rate +/- Inspiratory/expiratory control | Stabilize breath by breath variation | Periodic breathing Cheyne Stokes Respiration Treatment induced central sleep apnea Can worsen hypoventilation

## Slide 21: Servoventilation (SV)
Respironics BiPAP AutoSV
Targets 90-95% recent tidal volume
Mid-inspiration adjustment
EPAP adjustment for sustained periodic breathing or OSA
ResMed ASV
Targets 90% recent minute ventilation
Adjusts PS every 1/50th seconds
AutoASV adjusts EPAP for OSA
Hill Patent 2008/ 0000475
Apnea
ASV turned on
Patient Flow
ASV pressure

## Slide 22: Titrating ASV
Typically start EEP 4-5 unless severe obstruction
Slowly increase EEP for events or instability
Generally keep PS min at 3 and PS max at 15
  If difficulty tolerating pressure decrease PS max to 10
  If can’t eliminate centrals can try PS min of 0
Make sure NO mask leak
If hypoventilation switch to VAPS
If ASV doesn’t work try AutoSV and vice versa

## Slide 23: Volume Assured Pressure Support (VAPS)
Pressure Profile | Goal | Uses/ Cautions
Fixed or autoEPAP in response to obstruction Variable IPAP/PS targeting set goal volume or ventilation Backup rate Inspiratory/expiratory control | Stabilize breath by breath variation Ensure stable ventilation throughout entire night | Combination of periodic breathing and hypoventilation Supine or REM dominant disease Hypoventilation Neuromuscular

## Slide 24: iVAPS vs AVAPS
 | iVAPS | AVAPS
EPAP Adjustment | Fixed EPAP AutoEPAP (Astral) | SystemOne AVAPS- fixed EPAP Trilogy AVAPS- AE- set EPAPmin and max with REMstar AutoCPAP-like protocol
PS Adjustment | PS adjustments are made every 1/50th sec throughout inspiration to achieve goal alveolar ventilation with smooth transition. | Determines average PS provided over prior 2 min to achieve volume. If target volume differs from average recent ventilation (6 breaths), PS for next breath is changed at rate of 1/min if recent stable breathing and 0.5/min if unstable breathing. (AE mode allows maximum rate of pressure change from 1-5/min).

## Slide 25: Minute ventilation, alveolar ventilation and tidal volume
Tidal volume (BiPAP AutoSV and AVAPS)
  Volume to lung and anatomical dead space
  If ↑ respiratory rate, alveolar ventilation ↑
Minute Ventilation (ASV target)
  Respiratory rate X tidal volume
  If ↑ respiratory rate, alveolar ventilation ↓
Alveolar ventilation (iVAPS target)
  Minute ventilation – dead space ventilation
  If ↑ respiratory rate, then TV ↓ and alveolar ventilation is constant

## Slide 26: AVAPS-AE and iVAPS backup rate
AVAPS-AE backup rate
  Fixed or Autoset which is 2 breaths per min below average rate of recent 6 spontaneous breaths
iVAPS patient rate
  Patient rate drops to 2/3rd of set rate during spontaneous breathing and increases to set rate during period of apnea
From resmed.com

## Slide 27: Titrating VAPS
Start with EPAP at pressure preventing very severe obstructive events although lower may work
Start TV/ TVa per body weight chart
Start PS min 4 and PSmax of 15-20
Start RR/Patient Rate at respiratory rate
COPD has more dead space ventilation so expect to use higher Tva/TV
Start normal but try adding in ventilatory control settings in appropriate patients- ie, COPD, neuromuscular
Don’t just keep increasing EPAP and Tva/TV
Record tidal volumes and PS min and max achieved by patient throughout titration
Watch for setting increases that lead to lower patient TV for a given setting change

## Slide 28: Tip 6: Follow Tidal Volumes and IPAP max values when titrating iVAPS
Note: IPAP Max is maxed out for further increases of TVa at this point not significantly changing TV or lowering CO2
If increase patient rate, can increase ventilation

## Slide 29: Case 1: Hypoventilation/Hypoxia with Treatment emergent centrals on BPAP
BPAP 16/9 in REM
Stable Breathing
O2 Sat: 84%
TCCO2: 56
BPAP 17/9 in NREM
Treatment induced centrals
O2 Sat: 92%
TCCO2: 51

## Slide 30: Case 1: ASV treats treatment emergent centrals
ASV EEP 7:
O2 saturations improved to 93%
Breathing most stable in NREM
TCCO2 still 50
ASV EEP 8:
O2 saturations drop in REM to 88-89%
TCCO2 increases to 53
….but worsens hypoventilation

## Slide 31: Case 1: iVAPS improves centrals and hypoventilation
iVAPS EPAP 9 TVa 6 PR 20

## Slide 32: RAD Criteria for SV/VAPS/BPAP ST
To qualify for an E0471 Respiratory Assist Device (RAD)- BiPAP ST, ASV, iVAPS, AVAPS
  1. Central sleep apnea at baseline with CAHI >5 with OAHI < 5
  2. Treatment emergent centrals with CAHI > 5 with OAHI < 5 on BiPAP (If get centrals on CPAP, also get them on BiPAP before switching to backup rate machine)
    This must be present on at least 1 BiPAP pressure without higher pressures with either OAHI > 5 or higher pressures with overall AHI <5
  3. Hypoventilation Criteria on BiPAP- Requires ABG on BiPAP

## Slide 33: Ventilator Criteria For VAPS
To qualify for ventilator (Astral or Trilogy with iVAPS/AVAPS or BiPAP settings)
  Medical documentation of chronic respiratory failure
  ABG with PCO2 > 52 if COPD or obesity hypoventilation or PCO2>45 or MIP < 60 if Neuromuscular
  In this case do not need to meet either central criteria or 5 minutes below 88% on oxygen criteria
Expensive monthly rental fee
Must have persistent chronic respiratory failure to stay on
No compliance requirements to keep the machine

## Slide 34: Case 2: OSA/Hypoventilation/Hypoxia
69 yo male with DM, HTN, Morbid obesity (BMI 49), CAD s/p triple bipass admitted for CHF exacerbation in 2013
In 2005, dx with severe OSA and hypoventilation and didn’t tolerate CPAP 6
Split Study
Baseline on 4 L oxygen
  AHI 37
  Lowest O2 sat 77%
  Average ETCO2 52
  Peak ETCO2 89
Titration
  Lowest O2 sat 70%
  Peak ETCO2 99

## Slide 35: Case 2: PSG on 4 liters oxygen
What about the CO2 levels?

## Slide 36: Tip 7: Check ETCO2 waveform
ETCO2 correlates poorly with ABG in specific patient populations
  COPD
  Small lung volumes
  Obesity
  Acute respiratory illnesses
ETCO2 signal can be affected by
  Apnea: no airflow to the sensor
  Hypnopnea: if don’t reach end expiration
  Reduced tidal volume: don’t reach end expiration
  Obstruction: due to air trapping/lack of equilibration
  Mask leak: loss of signal
  High PAP flow: diluting signal
Noninvasive monitoring of CO2 during polysomnography: a review of the recent literature Christopher A. Gerdung, Adetayo Adeleye, and Valerie G. Kirk Curr Opin Pulm Med 2016, 22:527–534

## Slide 37: Case 2: BPAP 16/6 with 4 L O2
Why are saturations in 70s on BiPAP with 4L oxygen, when they were in 90s with 4L oxygen?

## Slide 38: Tip 8: Be aware of oxygen dilution with PAP
Oxygen flow diluted by room air flow (21% FiO2) from the machine
Higher pressure, more O2 flow needed for same FiO2
If the patient is on daytime O2, make sure they are switching to the appropriate (higher) level with PAP

## Slide 39: Case 2: BPAP ST 20/14 RR14 4 L O2
Tip: BPAP can cause squared airflow pattern that can be confused for obstruction
Overall:
Treatment emergent centrals
Prolonged centrals with BiPAP ST
Persistent Hypoxia
Persistent Hypoventilation
  REM hypoventilation with BPAP

## Slide 40: Next day AVAPS titration
  Next day comes back for AVAPS titration
    AVAPS with EPAP of 14, IPAP min of 18, IPAP max of 30, TV of 550 and 4 liters of oxygen and backup rate of 14 with a heated humidifier.

## Slide 41: Case 2: AVAPS EPAP 14 min TV 550 IPAPmin 18 IPAPmax 30 rate 10 O2 4L
Persistent central events
Oxygen Banding in NREM

## Slide 42: Case 2: AVAPS EPAP 14 min TV 550 IPAPmin 18 IPAPmax 30 rate 10 O2 4L
Persistent Hypoventilation worse in REM

## Slide 43: Case 2: AVAPS Followup
Uses AVAPS 4 L O2 nightly
  “Fights with it all night”
  Sleeps 4-5 hours on machine with awakenings after about 2 hours then has to take it off
Difficulty falling asleep and staying asleep, daytime fatigue, nocturia

## Slide 44: Case 2: iVAPS EPAP 10 TVa 5.6 PR 12 with 2 liter O2

## Slide 45: Case 2: Sleep Clinic Follow-up
Recommend iVAPS
  EPAP of 9 qnd TVa of 5.6
  PS min of 4 and PS max of 20
  Patient Rate of 14
  Ti min 0.3, Ti max 2, High Cycle
  2 liters oxygen
Took over 1 month to get approval to switch to iVAPS
“Great sleep on iVAPS”, ESS 5-> 0, sleeping 9:30-5 AM
Average use 6h57 min, 26/30 days over 4 hours,
Residual AHI 2.3

## Slide 46: Tip 9: Fixing hypoventilation helps hypoxia
Alveolar Gas Equation
  PAO2 = FiO2(Barometric Pressure – PH2O) – PACO2/Respiratory Quotient 760 47 0.8
  The more space CO2 takes up the less space for O2 in the alveola
    PAO2 = 0.21 (760 -47) – 40/0.8 = 99 mm Hg ~ > 95% SaO2
    PAO2 = 0.21 (760 -47) – 60/0.8 = 74 mmHg ~ 92% SaO2
    PAO2 = 0.21 (760 – 47) – 80/0.8 = 49 mmHg ~ 80% SaO2
Image from Becker Anesthesia Progress 2009

## Slide 47: 6 titration elements
Element | Exacerbating Factors
Obstruction | Worse REM, Supine
Hypoventilation | Worse REM/ Supine
Periodic breathing (PB)/Cheyne Stokes respiration (CSR) | Worse NREM/Supine Worse higher pressures Worse hypoxia/hypoventilation Worse backup rate
Irregular centrals (narcotic) | Narcotic effect could be worse at start of night
Insufficient effort (Neuro) | Worse REM/Supine
Baseline Hypoxia | Worse REM/Supine

## Slide 48: Titration hierarchy
Obstruction
CPAP
Hypoventilation
BPAP +/-respiratory control
Hypoxia
BPAP- w/ hypoventilation
Periodic Centrals
SV – w/o hypovent
VAPS- w hypovent
O2- w/o hypoventilation
VAPS

## Slide 49: Obesity Hypoventilation Titration
OSA
Hypoventilation only
No OSA
CPAP
Central events or Persistent ↑CO2 or Persistent ↓O2
BiPAP +/- respiratory control settings
Central criteria or ↑CO2 on max BiPAP
VAPS
VAPS + O2
BiPAP ST +/- O2
Suboptimal response to everything including lateral positioning
BiPAP/BiPAP ST/VAPS
  Long Ti
  Short rise time
May need large pressure support
Many need increased backup rate
Persistent ↓O2

## Slide 50: CHF Titration
Baseline CSA only
Normal O2
+/- ↑CO2
OSA +/- ↓O2
+/- ↑CO2
Hypoxia normal CO2
Hypoventilation +/- ↓O2
CPAP
BiPAP
Central criteria w/ nl CO2 but primarily CSA at baseline in patients with EF < 30%
Oxygen
BiPAP + O2
VAPS
VAPS + O2
Central criteria
BiPAP ST +/- O2
Persistent ↓O2
CPAP + O2
Suboptimal response to everything else including lateral positioning
Persistent ↓O2
Persistent ↓O2
Central criteria
w/ Nl CO2
SV if normal CO2 and not primarily CSA at baseline in patients with EF < 30%
Mixed CSA + OSA +/- ↓O2
+/- ↑CO2
Baseline CSA only
With ↓O2
Normal CO2
Persistent ↑CO2 +/- central criteria
Persistent CSA
Persistent Centrals
CPAP or SV with end expiratory rebreathing space +/- O2
Persistent Centrals with nl CO2
*Vary to meet insurance qualifications

## Slide 51: COPD Titration
OSA +/- hypoxia and +/- hypoventilation
Hypoxia without hypoventilation
Hypoventilation +/- Hypoxia
CPAP
Persistent ↑CO2 +/- central criteria
BiPAP
Central criteria or ↑CO2 on max BiPAP
Oxygen
Persistent ↓O2
w/ normal CO2
BiPAP + O2
VAPS
VAPS + O2
Central criteria
BiPAP ST + O2
Persistent ↓O2
CPAP + O2
Suboptimal response to everything else including lateral positioning
BiPAP/BiPAP ST/VAPS
  short Ti
  high cycle
  short rise time
Watch for EPAP or pressure support that start making CO2 rise more or lower TV as sign of worsening hyperinflation
Increase patient rate if hypoventilation despite recorded IPAP during titration at PSmax
Persistent ↓O2
Central criteria
Normal CO2
SV
*Vary to meet insurance qualifications
Persistent centrals

## Slide 52: Final Titration Tips
Beware of snore artifact and BPAP “squaring” leading to overtitration and missing central physiology
O2 may be able to be reduced once treat hypoventilation
VAPS and SV often need lower EPAP than BPAP
If one SV or VAPS device doesn’t work, try another
Higher pressures are not always better, but some evidence that COPD may require high PS

## Slide 53: Insurance Qualification Summary
Score central hypopneas to qualify for BPAP with backup (BPAP ST, SV, VAPS)
If VAPS best for controlling hypoventilation, but doesn’t qualify, consider if meets ventilator criteria
Almost impossible to switch from one Backup rate machine to other

## Slide 54: Final Titration Tips
Tech’s goal should be to show setting options in different stages and positions
MD may need to extrapolate optimal setting and follow adherence data residual AHI, oximetry and ABG to adjust after set-up
Think about the underlying physiology interacts with PAP
You may not be able to fix everything
Some patients can never be fixed in supine position

## Slide 55: Good luck titrating
Karin.johnson@baystatehealth.org
@drsleepykarin
For more information regarding how PAP modes work
  Johnson & Johnson Treatment of sleep disordered breathing with positive airway pressure devices: technology update. Medical Devices: Evidence and Research 2015:8: 425-437.