One-Step Strep A Rapid Test

## INTENDED USE

- One-Step Strep A Test is a lateral flow, one-step immunoassay for the rapid, qualitative detection of Group A Streptococcal antigen from throat swabs.

## SUMMARY

- Beta-hemolytic Group A Streptococcus is a major cause of upper respiratory infection such as tonsillitis, pharyngitis and scarlet fever. Early diagnosis and treatment of Group A Streptococcal pharyngitis has been shown to reduce the severity of symptoms and further complications, such as rheumatic fever and glomerulonephritis.

- Conventional methods used for the detection of the disease depend on the isolation and subsequent identification of the organism. These methods often require 24 - 48 hours to complete. Recent development of immunological techniques which can detect Group A Streptococcal antigen directly from throat swabs allow physicians to diagnose and administer therapy immediately.

## PRINCIPLE

The One-Step Strep A Test utilizes two site sandwich immunoassay technology for the detection of Group A Streptococcal antigen. The test device consists of plastic housing containing a membrane strip which has been precoated with rabbit anti-Strep A antibody on the test band region and goat anti-rabbit antibody on the control band region. A colored rabbit anti-Strep A polyclonal antibody-colloid gold conjugate pad is placed at the end of the membrane. During testing, the Strep A antigen is extracted from the throat swab using Extraction Reagents A, B and C. The extracted solution is then added to the sample well. The Strep A antigen reacts with colored antibody-colloidal gold conjugate to form Strep A antigen-antibody complexes. The mixture then moves chromatographically across the membrane to the immobilized rabbit anti-Strep A antibody at the test band region. If Strep A antigen is present in the specimen, a pink colored sandwich of solid phase/Strep A antigen/gold conjugate is formed on the test band region. Absence of the pink colored band at the test band region indicates a negative result. Regardless of the presence of Strep A antigen, as the extracted mixture continues to move laterally across the membrane to the immobilized goat anti-rabbit antibody test band region, a pink colored band at the control region will always appear. The presence of this colored band serves as: 1) verification that sufficient volume has been added, 2) verification that proper flow is abtained, and 3) reagent control.

## REAGENTS AND MATERIALS SUPPLIED

- Extraction Reagent A: NaOH (5ml).

- Extraction Reagent B: Acid (5ml).

- Extraction Reagent C: Oxidation Reagent (5ml).

- Test Device: 20.

- Sterile Throat Swabs: 20

## MATERIALS REQUIRED BUT NOT SUPPLIED

- Specimen collection container

- Timer

## STORAGE AND STABILITY

- Store as packaged in the sealed pouch at 2-30°C.

- The test device must remain in the sealed pouch until use.

- DO NOT FREEZE.

- Do not use beyond the expiration date.

## PRECAUTIONS

- For professional and IN VITRO diagnostic use only.

- The test device should remain in the sealed pouch until use. Do not use after the expiration date.

- Extraction Reagent B is slightly caustic. Avoid contact with eyes or mucous membranes. In the event of accidental contact, wash thoroughly with large volumes of water.

- Standard guidelines for handing infectious agents and chemical reagents should be observed throughout all procedures. All contaminated waste such as swabs, Strep A, test devices and extracts should be properly disposed.

- To obtain accurate results, package insert instruction must be followed.

## SPECIMEN COLLECTION AND STORAGE

Collect throat swab specimens by standard clinical methods such as those described by Facklam (1) and Ross (5). Use only Dacron or Rayon tipped sterile swabs with plastic shafts such as those provided. Do not use calcium alginate, cotton tipped, or wooden shafted swabs.

It is recommended that swabs specimens be processed as soon as possible after collection. If swabs are not processed immediately, they should be placed into a sterile, dry, tightly capped tube or bottle and refrigerated. If a liquid transport methods is employed, use Modified Stuart’s Transport Media as outlined in the manufacture’s instructions. Do not use transport media formulas including charcoal or agar. Swabs can be stored at room temperature ( 15-30℃) up to 4 hours, or refrigerated (2-8℃) up to 24 hours.

If a bacteria culture is desired, lightly streak the swab on a 5% sheep blood agar plate before using it in the SureStep Strep A Test. Extraction reagents kill the bacteria on swabs and make them impossible to culture. Alternatively, a subsequent second swab sample may be taken for culture procedure.

## TEST PROCEDURE

Procedural Notes

- If specimen swabs or any One-Step Strep A Test reagents have been refrigerated, allow them to equilibrate to room temperature before testing.

- Label the device with patient or control identification. To avoid cross contamination, do not allow the tips of the reagent bottles to come in contact with sample swabs and Extraction Tubes.

Extraction Procedure

- Place an Extraction Tube in the designated areas of the Workstation. Add 4 drops of Extraction Reagent A to the Extraction Tubes.

- Add 4 drops of Extraction Reagent B to the tube. Immediately place the throat swab specimen in the tube. Use a circular motion to roll the swab against the side of the Extraction Tube so that the liquid is expressed from the swab and reabsorbed. Let stand for a minimum of 1 minute at room temperature and maximum of 15 minutes.

- Add 4 drops of Extraction Reagent C to the tube. Squeeze the swab firmly against the tube to expel as much liquid as possible from the swab. Discard the swab.

## Test Procedure

- Cap the Extraction Tube with the attached dropper. Add 3 drops of all the extracted solution to the sample well of the test device.

- Read result in 5 minutes. Depending on the number of organisms on the swab, a positive result may be visible as soon as 1 minute. However, to confirm a negative result the complete reaction time of 5 minutes is required. Do not read result after 10 minutes.

## INTERPRETATION OF RESULTS

|  | Positive Two colored lines should be observed in the viewing region. The line in the test region (T) is the probe line; the line in the control region (C ) is the control line, which is used to indicate proper performance of the test. The color intensity of the test line may be weaker or stronger than that of the control line. |
| --- | --- |
|  | ### Negative The control line appears in the test window, but the test line is not visible. |
|  | Invalid No line appears in the control region. Under no circumstances should a positive sample be identified until the control line forms in the viewing area. If the control line does not form, the test result is inconclusive and the assay should be repeated. |

## LIMITATION OF PROCEDURE

- The accuracy of the test dependents on the quality of the swab sample. False negative may result from improper sample collection or storage. A negative result may be obtained from patients at the onset of the disease due to low antigen concentration. Therefore, when a patient suspected of having infection, additional testing using the culture method is required.

- The test does not differentiate asymptomatic carriers of Group A Streptococcus from those with infection. If clinical signs and symptoms are not consistent with laboratory test results, a follow-up throat culture is recommended.

- In rare cases, test specimens heavily colonized with Staphylococcus auras can yield false positive results. If clinical signs and symptoms are not consistent with clinical test results, a follow-up culture procedure should be performed.

- Respiratory infections, including pharyngitis, can be caused by Streptococci from serogroups other than Group A, as well as by other pathogens.

- As in the case of any diagnostic procedure, the results obtained with this test should be used in conjunction with other information available to the physician.

## QUALITY CONTROL

- Good laboratory practice recommends the use of control materials to ensure proper kit performance. A positive control containing heat-killed Group A Streptococcus and a negative control containing heat-killed group B Streptococcus are sale separately by manufacture.

- Positive control: Add 4 drops of Extraction Reagent A and 4 drops of Extraction Reagent B to an Extraction Tube. Thoroughly mix the Positive Control by shaking the bottle vigorously. Add 1 drop of Positive Control to the tube. Place a sterile swab into the tube and swirl. Continue Extraction Procedure Step 3, followed by Test Procedure.

- Negative Control: Add 4 drops of Extraction Reagent A and B to an Extraction Tub. Thoroughly mix the Negative Control by shaking the bottle vigorously. Add 1 drop of Negative Control to the tube. Place a sterile swab into the tube and swirl. Continue Extraction Procedure Step 3, followed by the Test Procedure.

## PERFORMANCE CHARACTERISTICS

##### Sensitivity Study

- To determine the analytical sensitivity of the One-Step Strep A Test, Group A Streptococcus bacteria were grown in broth culture. The detection limit of the One-Step Strep A Test was determined to be 2.5 × 105 organisms per test.

##### Specificity Study

- To determine the specificity of the One-Step Strep A Test to Group A Streptococcal bacteria, the following Group A Streptococcal Strains at different levels of organisms per test were examined. Positive results obtained at the level of 2.5 × 105 organisms/test for all Strep A strains indicate that One-Step Strep A Test is sensitive to all Group A Streptococcal bacteria.

- Group A Streptococcal Strains:

- SS-091 	    SS-410		SS-492		SS-496	    SS-633	     SS-634     SS-635        SS-721		SS-754		SS-799        ATCC-19615

Cross-reactivity studies with organisms likely to be found in the respiratory tract were also performed using the One-Step Strep A Test. The following organisms were tested at 1 × 108 organisms/test.

Group B Streptococcus 	Group C Streptococcus

Group D Streptococcus 	Group F Streptococcus

Group G Streptococcus 	Pseudomonas aeruginosa

Streptococcus bovis		Staphylococcus aureus

Proteus vulgaris		     Streptococcus faecalis

Staphylococcus epidermidis	 Escherichia coli

Staphylococcus saprophyticus	Streptococcus mitis

Corynebacterium diphtheriae  Neisseria lactima

Neisseria gonorrhoeae	    Streptococcus pneumoniae

Streptococcus mutans	     Moraxella catarrhalis

Bordetella pertussis	     Neisseria sicca

Neisseria meningitidis	     Streptococcus sanguis

Streptococcus salivarius	Candida albicans

Neisseria subflava	   Haemophilus parahaemolyticus

- Negative results in all above cases indicate that the One-Step Strep A Test is specific to Strep A bacteria only.

## Correlation Study

- A correlation study of the One-Step Strep A Test and the conventional culture tests has been determined in multi-center clinical evaluations. Throat swab specimens were taken from patients exhibiting symptoms of pharyngitis. The swabs were then used to inoculate blood agar plates prior to testing with the One-Step Strep A Test. Beta-hemolytic colonies from the blood agar plates were confirmed as Group A Streptococcus using serologic streptococcal grouping methods. Strep A was reported as present or not present.

- The qualitative results are summarized as follows:

| Strep A |  | + | - | Total |
| --- | --- | --- | --- | --- |
|  | + | 21 | 2 | 23 |
| Culture | - | 3 | 35 | 38 |
|  | Total | 24 | 37 | 61 |

- Sensitivity:        91.3%

- Specificity:        92.1%

- Over all accuracy:   91.8%

## Physician Office Laboratory Studies

An evaluation of One-Step Strep A Test was conducted at three physician Office Laboratory sites, using a panel of coded samples containing Negative Control, Low positive, Medium Positive and High Positive specimens. One hundred percent (100%) agreement to the expected results was obtained.

ATLAS LINK, INC

No. 811 Zeyang Plaza, No.166 Fushi Road

Shijingshan Dist., Beijing 100043, China

Tel: 86-10-88909112  Fax: 86-10-88909115

CIRIANO GLOBAL S.L. (UNILATEX GROUP)

CIF: B50927532 C/Blancas 4-6, 1 B Oficina 1

50001 Zaragoza Spain

Tel: 34-976910399

Latest revision 09.05.2011

USA Headquarters:

10758-A Ambassador Drive

MANASSAS, VA 20109 USA

Tel: 703-3663851