Quicking Biotech Co., Ltd.

No.1998, South Yanggao Road, Shanghai, China. Tel: 86-21-5094 7188

CA-125 RAPID TEST

COMMON NAME

CA-125 Rapid Test Kit

PACKING SIZE

1 cassette / pouch, 20 pouches / box

INTENDED USE

For the qualitative detection of the ovarian cancer antigen CA-125 concentration in human serum.

PRINCIPLE

Sandwich method immunochromatographic assay

KIT COMPONENTS

20 pouches of CA-125 cassette

One product insert

STORAGE AND STABILITY

Store at room temperature (2-30°C). Don’t freeze. Shelf life is 18 months.

SAMPLE REQUIREMENT

Clean serum, undiluted. If precipitate exists in the sample, please do centrifugation before use.

TEST PROCEDURE

1) Reading carefully the instructions before use.

2) Allow all kit components and specimen to reach room temperature prior to testing.

3) Remove the test device from the foil pouch prior to use.

4) Place the test device horizontally, Drop 3 drops of the specimen into “S” hole.

5) Interpret test results in 10-15 minutes. Do not interpret after 20 minutes. 40 U/mL and above is considered as positive result.

INTERPRETATION OF THE RESULT

- Positive: The presence of two bands in both C zone and T zone.

- Negative: only one band appears in C zone.

- Invalid: No visible band appear in C zone, no matter if there is band in T zone. Please repeat the testing for a valid result.

PRECAUTIONS

1) For best results, strict adherence to these instructions is required.

2) Do not open or remove test kit from their individually sealed pouches until immediately before their use.

3) Do not reuse test kit.

4) All reagents must be at room temperature before running the assay.

5) Do not use reagents beyond the stated expiration date marked on the package label.

LIMITATION

1) This Test is for in-vitro diagnostic use only.

2) For more accuracy of diagnosis, additional follow-up testing using other medical methods is recommended.

Quicking Biotech Co., Ltd.

No. 1998, South Yanggao Road,

Shanghai, 200125 China.

Tel: +86-21-5094 7188

Fax: +86-21-5094 7186

Website: www.quicking.cn