##### Drug Screen Card

##### One-Step Panel Test

For in vitro Diagnosis Use

## INTENDED USE

- The Drug Screen card is a lateral flow, one-step panel immunoassay for the qualitative detection of multiple drug metabolites in human urine at the following cut-off concentrations.

Compound                        cut-off

MOR                           300ng/ml

PCP                            25ng/ml

MAMP                        1000ng/ml

COC 						   300ng/ml

THC						    50ng/ml

AMP                          1000ng/ml

BZO                            300ng/ml

BAR                            300ng/ml

MDMA                         500ng/ml

MTD                           300ng/ml

- This product is used to obtain a visual, qualitative result and is intended for use by drug testing professionals in drug testing programs. The assay should not be used without proper supervision and is not intended for over the counter sale to lay persons.

- This assay provides only a preliminary analytical test result. A more specific alternative chemical method must be used in order to obtain a confirmed analytical result. Gas chromatography/mass spectrometry (GC/MS) has been established as the preferred confirmatory method by the National Institute on Drug Abuse (NIDA). Clinical consideration and professional judgment should be applied to any drug of abuse test result, particularly when preliminary positive results are indicated.

## SUMMARY

- Urine based screened tests for drugs of abuse of abuse range from simple immunoassay tests to complex analytical procedures. The speed of immunoassays has made them the most widely accepted method for screening urine for drugs of abuse.

- Our Drug screen card is based on the principle of highly specific immunochemical reactions of antibodies, which are used for the analysis of specific compounds in biological fluids. The drug screen card is rapid, visual, competitive panel immunoassays that can be used for the simultaneous, qualitative detection of multiple drug metabolites in human urine.

- Sample adulteration is a serious problem in forensic urine sample and to exclude liquid substitutions and grossly diluted urine samples.

## PRINCIPLE

- The Drug Screen Card is a one-step Panel Test immunoassay in which a chemically labeled drug (drug conjugate) competes with the drug, which may be present in urine for limited antibody binding sites. The test device contains a membrane strip, which was pre-coated with drug conjugate on the test band region. The colored monoclonal antibodies-colloidal gold conjugate pad is placed at the end of the membrane. The colored antibody-colloidal gold conjugate moves along with urine, chromatographically by the capillary action, across the membrane, in the absence of drug in the urine, the colored antibody colloidal gold conjugate attaches to the drug conjugate on the test band region to form a visible line as the antibody/drug conjugate complexes. Therefore, the formation of a visible precipitant in the test band region occurs when the test urine is negative for the drug. When the drug is present in the urine, the drug/metabolite antigen competes with the drug conjugate on the test band region for the limited antibody sites. When a sufficient amount of drug is present, it will fill the limited antibody binding sites. This will prevent attachment of the colored antibody-colloidal gold conjugate to the drug conjugate zone on the test band region. Therefore, absence of the color band on the test region indicates a positive result.

- A control or reference band with a different antigen/antibody reaction is also added to the immunoassay membrane strip to indicate that the test is performed properly. This control line should always appear regardless of the presence of drug or metabolite. This means that negative urine will produce two colored bands, and positive urine will produce only one band. The presence of this colored band in the control region also serves as verification that 1) sufficient volume has been added, and 2) that proper flow was obtained.

## REAGENTS AND MATERIALS SUPPLIED

- Individually wrapped test devices.

- One instruction sheet.

## MATERIALS  REQUIRED BUT NOT SUPPLIED

- Specimen collection container

- Timer

## STORAGE AND STABILITY

- Store as packaged in the sealed pouch at 2-30°C.

- The test device must remain in the sealed pouch until use.

- DO NOT FREEZE.

- Do not use beyond the expiration date.

## PRECAUTIONS

- For professional and IN VITRO diagnostic use only.

- The test device should remain in the sealed pouch until use. Do not use after the expiration date.

- Urine specimens may be potentially infectious. Proper handing and disposal methods should be established.

- Avoid cross-contamination of urine samples by using a new specimen collection container for each urine sample.

## SPECIMEN COLLECTION AND STORAGE

- The Drug Screen Card is formulated for using with urine specimens. Fresh urine does not require any special handing or pretreatment.

- Testing should be performed immediately after the specimens have been collected. Do not leave the specimens at room temperature for prolonged periods. Specimens may be stored at 2-8°C for up to 2 days.  For long-term storage, specimens should be kept below -20°C.

- Specimens that have been refrigerated must be equilibrated to room temperature prior to testing. Specimens previously frozen must be thawed, equilibrated to room temperature, and mixed thoroughly prior to testing.

- Note: Urine specimens and all materials coming in contact with them should be handled and disposed of as if capable of transmitting infection. Avoid contact with skin by wearing gloves and proper laboratory attire.

## TEST PROCEDURE

- Remove the test card from the protective pouch (bring the device to room temperature before opening the pouch to avoid condensation of moisture on the membrane).

- In the spaces provided for identification on the test card, label with patient or control identifications, operator and date. Remove the cap to expose the absorbent tips.

- Immerse the sample strips into the urine specimen until it is observed that a pink color begins to migrate across each result window (approximately 30seconds). Do not allow urine to touch the plastic device when immersing the sample strips. Replace the cap to cover the sample strips. Then place the test card on a flat surface.

- Read results within 3 to 8 minutes. Do not read results after 8 minutes.

## INTERPRETATION OF RESULTS

|  | Negative: Two colored lines appear adjacent to each other in the result window. This indicates that no drug above the cut-off level has been detected. The test line intensity may be any shade of pink and may be weaker or stronger than that of the control line. All tests are independent of each other, it is not recommended to compare results or color intensity of the test lines. |
| --- | --- |
|  | Positive: Only one colored line appears in the control region (C). No test line appears in the test region (T). A positive result should not be considered conclusive and should be confirmed with a more specific alternative chemical method such as GC/MS. |
|  | Invalid: If no control line appears in the result window, the test should be considered void. Improper testing procedures, sample tampering or test deterioration probably occurred. Re-test with a new card and if condition persists, contact Technical Assistance. |

- Note: Very faint lines on the test region indicates that the drugs in the sample is near the cut-off level for the test. These samples should be re-tested or confirmed with a more specific method before a positive determination is made.

## LIMITATION OF PROCEDURE

- The assay is designed for human urine use only.

- A positive result with any of the tests indicates the presence of a drug/metabolite only and does not indicate or measure intoxication.

- There is a possibility that technical or procedural errors as well other substances or factors not listed may interfere with the test and cause false results. See SPECIFICITY for lists of substances that will produce positive results, or that do not interfere with test performance.

- If it is suspected that the samples have been mislabeled or tampered with, a new specimen should be collected and the test should be repeated.

## QUALITY CONTROL

- Good laboratory practice recommends the use of control materials to ensure proper kit performance. Quality control specimens are available from commercial sources. When testing the positive and negative controls, use the same assay procedure as with a urine specimen.

## PERFORMANCE CHARACTERISTICS

##### Accuracy

##### The accuracy of Drug Screen Card was evaluated in comparison to a commercially available immunoassay. One hundred twenty two (122) urine samples, collected from presumed non-user volunteers, have been tested by both procedures with 100% agreement.

##### In a separate study, fifty (50) urine samples, obtained from a clinical laboratory where they were screened and confirmed as positives by the commercially available immunoassay and GC/MS, were tested with the Drug Screen Card.

#### Precision

The precision of the Drug Screen Card was determined by conducting the test with spiked blind controls. The control at a concentration of 50% below the cut-off level will give a negative result and the control at a concentration of 200% above the cut-off level will give a positive result.

#### Specificity

The specificity for the Drug Screen Card was tested by adding various drugs, drug metabolites, and other compounds that are likely to be present in urine. All compounds were prepared in drug-free normal human urine.

The following structurally related compounds produced positive results when tested at levels equal to or greater than the concentrations listed below.

| Compound | Concentration (ng/ml) |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Phencyclidine |  |  |  |  |  |
| Tenocyclidine | 2,000 |  |  |  |  |
| Morphine |  |  |  |  |  |
| Morphine |  |  | 2000 |  |  |
| Codeine |  |  | 2000 |  |  |
| Ethyl Morphine |  |  | 2000 |  |  |
| Hydrocodone |  |  | 25,000 |  |  |
| Hydromorphone |  |  | 25,000 |  |  |
| Morphine-3---d-glucuronide |  |  | 5000 |  |  |
| Thebaine |  |  | 75,000 |  |  |
| Methamphetamine |  |  |  |  |  |
| (+/-)3,4-Methamphetamine |  |  |  |  | 1000 |
| D-Amphetamine |  |  |  |  | 50,000 |
| Chloroquine |  |  |  |  | 50,000 |
| (+/-)-Ephedrine |  |  |  |  | 50,000 |
| (-)-Methamphetamine |  |  |  |  | 25,000 |
| (+/-)3,4-methylenedioxymethampheta- Mphetamine(MDMA) |  |  |  |  | 2000 |
| -Phemylethylamine |  |  |  |  | 50,000 |
| Ranitidine |  |  |  |  | 50,000 |
| Trimethobenzamide |  |  |  |  | 10,000 |
| Cocaine |  |  |  |  |  |
| Benzoylecogonine |  | 300 |  |  |  |
| Cocaine |  | 300 |  |  |  |
| THC |  |  |  |  |  |
| 11-nor-9-THC-9-COOH |  |  |  | 50 |  |
| 11-nor-8-THC-9-COOH |  |  |  | 50 |  |
| 11-hydroxy-9-Tetrahydrocannabinol |  |  |  | 2,500 |  |
| 8-Tetrahydrocannabinol |  |  |  | 7,500 |  |
| 9-Tetrahydrocannabinol |  |  |  | 10,000 |  |
| Cannabinol |  |  |  | 10,000 |  |
| Cannabidiol |  |  |  | 100,000 |  |

Amphetamine

| d-Amphetamine | 1,000 |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| l-Amphetamine | 10,000 |  |  |  |  |
| (+/-)3,4-methylenedioxy- amphetamine | 5,000 |  |  |  |  |
| phentemine | 1,000 |  |  |  |  |
| Tenocyclidine | 75,000 |  |  |  |  |
| Barbiturates |  |  |  |  |  |
| Secobarbital |  | 300 |  |  |  |
| Allobarbital |  | 1000 |  |  |  |
| Amobarbital |  | 1000 |  |  |  |
| Aprobarbital |  | 300 |  |  |  |
| Barbital |  | 300 |  |  |  |
| Butabarbital |  | 300 |  |  |  |
| Butalbital |  | 2000 |  |  |  |
| Butethat |  | 300 |  |  |  |
| Pentobarbital |  | 300 |  |  |  |
| Phenobarbital |  | 300 |  |  |  |
| Benzodiazepines |  |  |  |  |  |
| Oxazepam |  |  | 300 |  |  |
| Alprazolam |  |  | 150 |  |  |
| Bromazepam |  |  | 800 |  |  |
| Chlordiazepoxide |  |  | 300 |  |  |
| Clobazam |  |  | 200 |  |  |
| Clonazepam |  |  | 25,000 |  |  |
| Chlorazepam |  |  | 100 |  |  |
| Delorazepam |  |  | 6,000 |  |  |
| Diazepam |  |  | 150 |  |  |
| Estazolam |  |  | 300 |  |  |
| Flurazepam |  |  | 300 |  |  |
| Lorazepam1 |  |  | 500 |  |  |
| Lormetazepam |  |  | 1,000 |  |  |
| Medazepam |  |  | 2,000 |  |  |
| Nitrozepam |  |  | 1,000 |  |  |
| Nordiazepam |  |  | 100 |  |  |
| Prazepam |  |  | 1,000 |  |  |
| Temazepam |  |  | 150 |  |  |
| Triazolam |  |  | 1500 |  |  |

# MDMA

MDMA                            500

MDA                             2000

Cannabidiol                  100,000

MTD

2-ethylidene-1,5-dimethyl-3, 3-diphenylpyolidine                50,000

Methadol                         1,000                         	1,000