OneStep IgE InstaTest

Cat. No. 15150-50  (50 tests)

# introduction

InstaTest IgE is an immunochromatographic in vitro assay for detection of elevated levels of immunoglobulin E in human serum or plasma.

# Elevated concentration of IgE is commonly associated with allergic pathologies. It may also be caused by autoimmune diseases and some infections.

# Rapid detection of IgE using InstaTest IgE will help in diagnosis and monitoring of IgE-mediated disorders.

InstaTest IgE detects IGE concentrations of 80 IU/ml and higher. The test kit allows performing 50 tests. It contains 50 sealed foil pouches, with one immunochromatographic testing device and one disposable sample dropper in each pouch. A testing device is a test membrane with anti-IGE antibodies and conjugates in a stabilizing matrix containing proteins and sodium azide, enclosed in plastic casing.

The result is read visually by observing the symbols developing in the test ("T") and control ("C") zones of the result window. Distinctive control line serves as an indicator of proper performance of the test device.

InstaTest IgE is for in vitro diagnostic use only.

Storage conditions

Store below 28C; do not freeze.

TEST PROCEDURE AND INTERPRETATION OF RESULTS

Sample collection

Collect blood using a proper aseptical technique. Allow the blood to clot completely. Centrifuge to obtain clear serum. InstaTest IgE works also on human plasma. If necessary, the sample may be stored in a refrigerator for up to 48 hours, or frozen for prolonged storage.

For each test, approximately 0.15 ml of serum or plasma is required.

NOTES: Do not open the foil pouch until you are ready to perform the test.

If a sample or a test kit was refrigerated, bring unopened pouch

and test sample to room temperature before performing the test.

Test procedure

- Open a foil pouch by tearing along the "notch". Remove the testing device and the sample dropper. Place the testing device on a flat horizontal surface.

- Holding the dropper vertically, carefully dispense exactly four drops of the serum or plasma into the sample well (marked "S") of the testing device.

3.   Read the test results at five minutes.

Interpretation of the test results

The result is indicated by the two color lines (bands): the control line in the  control ("C") region of the test membrane, and the test line in the test ("T") region of the membrane.

|  | Positive Two pink-rose lines (bands) are visible in the control ("C") and test ("T") areas of the test window. The intensity of the test line may be less than that of the control line; this still means positive result: IgE concentration in the sample is 80 IU/ml or more. |
| --- | --- |
|  | Negative The control line appears in the test window, but the test line is not visible. IGE level in the test sample is below the InstaTest IgE sensitivity of 80 IU/ml. |
|  | Invalid The test is invalid if the control line is not visible at five minutes. The test failed, or the test procedure was not followed properly. Verify the test procedure and repeat the test with a new testing device. |

NOTE: Do not interpret InstaTest IgE results after more than five minutes. To avoid confusion, discard the testing device after registering the test result at five minutes.

Rev. 4/99