HAV IgM Gold Rapid Screen Test

【For the qualitative detection of Hepatitis A Virus IgM antibody in serum /plasma and whole blood 】

# INTENDED USE

The HAV IgM Rapid Test is a Chromatographic immunoassay (CIA) for direct qualitative detection of Hepatitis A Virus IgM antibody (HAV) in human serum or plasma and whole blood.

# PRINCIPLE

The HAV IgM RST is a chromatographic immunoassay (CIA) for the detection of IgM antibody to Hepatitis A Virus in human serum or plasma and whole blood. Specific antigen against HAV is pre-coated onto membrane as a capture reagent on the test region. During the test, specimen is allowed to react with the colloidal gold particles, which have been labeled with other specific antibodies. If HAV IgM antibody present, a pink colored band will develop on the membrane in proportion to the amount of HAV IgM antibody presented in the specimen. Absence of this pink colored band in the test region suggests a negative result. To serve as a procedural control, a pink colored band in the control region will always appear regardless the presence of HAV IgM antibody.

# STORAGE

The kits should be stored at temperature 4-28°C, the sealed pouch for the duration of the shelf life (24 months).

# SPECIMEN COLLECTION

Plasma

- Have a certified phlebotomist collect whole blood into a purple, blue or green top collection tube (containing EDTA, citrate or heparin, respectively) by veinpuncture.

- Separate the plasma by centrifugation.

- Label and store the plasma at 2-8°C for up to two weeks. Plasma may be frozen at -20°C for up to one year.

Serum

- Have a certified phlebotomist collect whole blood into a red top collection tube (containing no anticoagulants) by veinpuncture.

- Allow the blood to clot.

- Separate the serum by centrifugation.

- Label and store the plasma at 2-8°C for up to two weeks. Serum may be frozen at -20°C for up to one year.

# PROCAUTION

1.  FOR IN VITRO DIAGNOSTIC USES ONLY

2.  All patient samples should be treated as if capable of transmitting diseases.

3.  Do not interchange reagents from different lots or use test kit beyond expiration date.

4. Icteric, lipemic, hemolysed, heat treated and contaminated sera may cause erroneous results.

REAGENTS AND MATERIALS PROVIDED

1.  One pouched cassette with desiccant.

2.  One piece of operating instruction.

PROCEDURE

1. Mix 10ul serum/plasma and whole blood with 10ml salty water (0.85% NaCL). (That means 1:1000 diluent)

2. Add 70-100 ul or 2-3 drops of diluent specimen into sample well. Observe the result in 15 minutes.

3. Do not read the result after 15 minutes.

# INTERPRETATION OF RESULTS

Negative: Only one colored band appears on the control (C) region.  No apparent band on the test (T) region.

This indicates that no HAV IgM antibody has been detected.

Positive: In addition to a pink coloured control (C) band, a distinct pink colored band will also appear in the test (T) region. This indicates that the specimen contains HAV IgM antibody.

Invalid: A total absence of color in both regions is an indication of procedure error and/or the test reagent has deteriorated. The test should be repeated using a new device.

Negative           Positive           Invalid

# LIMITATIONS

1.  The test is to be used for the qualitative detection of HAV IgM antibody.

2.  A negative result does not rule out infection by Hepatitis A because the HAV IgM antibody may not be present in sufficient quality.

Latest revision 09.05.2011

ATLAS LINK, INC

No. 811 Zeyang Plaza, No.166 Fushi Road

Shijingshan Dist., Beijing 100043, China

Tel: 86-10-88909112  Fax: 86-10-88909115

USA Headquarters:

10758-A Ambassador Drive

MANASSAS, VA 20109 USA

Tel: 703-3663851

CIRIANO GLOBAL S.L. (UNILATEX GROUP)

CIF: B50927532 C/Blancas 4-6, 1 B Oficina 1

50001 Zaragoza Spain

Tel: 34-976910399