THE TEST

One Step LH Urine Ovulation Test is fast and easy-to-use. It is a qualitative test that can predict when there is a LH (luteinizing hormone) surge, and in turn when you are likely to ovulate.

| Your Cycle Length | Start To Test On |
| --- | --- |
| 21 days | Day 6 |
| 22 days | Day 6 |
| 23 days | Day 7 |
| 24 days | Day 7 |
| 25 days | Day 8 |
| 26 days | Day 9 |
| 27 days | Day 10 |
| 28 days | Day 11 |
| 29 days | Day 12 |
| 30 days | Day 13 |
| 31 days | Day 14 |
| 32 days | Day 15 |
| 33 days | Day 16 |
| 34 days | Day 17 |
| 35 days | Day 18 |
| 36 days | Day 19 |
| 37 days | Day 20 |
| 38 days | Day 21 |

WHEN TO BEGIN TESTING

First, you must determine the length of your menstrual cycle.  This is the number of days from the first day of your menstrual bleeding to the day before your next bleeding begins again.  Please refer to the chart to determine when you should start testing. If your cycle is shorter than 21 days or longer than 38 days, consult your doctor. If you do not know your cycle length, you may begin the test 11 days after your first period since the average cycle length is 28 days. Perform 1 test each day over a 5 day period or until the LH surge has been detected.

URINE COLLECTION

- Do not use first morning urine samples as LH is synthesized in your body early in the morning.  It will not show up in your urine until later in the day.

- The best time to collect your urine is between 10am - 8pm.

- Collect urine at about the same time each day.

- Reduce liquid intake about 2 hours before collecting your urine as a diluted urine sample can prevent the test from detecting LH surge.

TEST PROCEDURE

- Allow the sealed LH Cassette test pouch, patient’s urine sample and control solution to reach room temperature (18-30C).

- Remove the LH Cassette test device from the sealed pouch.

- Draw urine sample to the urine dropper provided.

- Hold the urine dropper upright about 1/2 inch above the urine well of the device.  Add 5 drops of urine to this well.

- Lay the test device on a flat surface with the result window facing up. Wait for colored bands to appear. Depending on the concentration of LH in the test specimen, positive results may be observed in as short as 40 seconds. However, to confirm negative results, the complete reaction time (10 minutes) is required. Do not read results after 30 minutes.

INTERPRETATION OF RESULTS

- No LH Surge: Only one color band appears on the control (C) region or the test (T) band is lighter than the control band.  There is no LH surge.

- LH Surge:  If two color bands are visible and the test (T) band is equal to or darker than the control (C) band, one will probably ovulate in the next 24-48 hours.  If one wants to be pregnant, the best time to have intercourse is after 24 but before 48 hours.

- Invalid: No visible band at all.  Repeat test with a new test kit.

Latest revision 09.05.2011

ATLAS LINK, INC                                                   USA Headquarters:

No. 811 Zeyang Plaza, No.166 Fushi Road             10758-A Ambassador Drive

Shijingshan Dist., Beijing 100043, China                MANASSAS, VA 20109 USA

Tel: 86-10-88909112  Fax: 86-10-88909115           Tel: 703-3663851

CIRIANO GLOBAL S.L. (UNILATEX GROUP)

CIF: B50927532 C/Blancas 4-6, 1 B Oficina 1

50001 Zaragoza Spain

Tel: 34-976910399